## Slide 1: REPROCESSING OF SURGICAL INSTRUMENTS AND EQUIPMENT & WASTE MANAGEMENT

## Slide 2: INRODUCTION
This is the making of re-usable instruments and equipment infection free and safe for use by the patient/client without transfer/transmission of infections and diseases by the following methods: decontamination, cleaning, and either sterilization or high-level disinfections (HLD)

## Slide 3: Decontamination
First step in handling used instruments and equipment
The recommended decontamination agent is a 0.5 percent chlorine solution.
Change solution every morning, or after 8 hours, or more often if the solution becomes visibly dirty
A 0.5 percent chlorine solution can be made from readily available liquid chlorine or chlorine tablets (NaDCC) using this formular:

## Slide 4
Total Parts (TP) water = (percentage chlorine in manufacturers concentration ÷ % desired chlorine concentration) – 1
E.g. To make a 0.5% chlorine solution from 5 percent concentrated chlorine solution:
TP water: (5.0% ÷ 0.5%) - 1= 10 - 1 = 9
It means, add 1 part concentrated solution to 9 parts water.
Cover containers containing 0.5 percent chlorine solution and protect them from light.
Note: Do not mix chlorine solutions with ammonia-based solutions, because toxic gas might be produced

## Slide 5
Decontaminate large surfaces that might have come in contact with blood and body fluid by wiping them with a cloth soaked in the 0.5 percent chlorine solution
Decontaminating Used Instruments
Decontaminate the instruments while wearing gloves
Immediately after use, place all instruments in 0.5% chlorine solution, for 10 minutes to inactivate most organisms, including HBV and HIV.
Use plastic, noncorrosive containers for decontamination to prevent sharp instruments from getting dull and/or rusted

## Slide 6
Remove instruments from chlorine solution after 10 minutes and immediately rinse them with cool water to remove residual chlorine before thoroughly cleaning them.
Remove gloves and dispose of them appropriately.
Note: To prevent rusting, do not soak metal instruments in water for more than one hour, even if they are electroplated.

## Slide 7: Cleaning
After decontamination and prior to disinfecting or sterilizing, all instruments and equipment MUST be cleaned to remove organic materials or chemical residue.
This prevents chemical interaction between the organic matter and the disinfecting/sterilizing chemicals preventing effective sterilization/disinfection
Use liquid soap or enzymatic detergent. Liquid soap suspends grease, oil, and other foreign matters in solution so that they can be removed easily by the cleaning process. Do not use abrasive cleaner e.g. steel wool to prevent scratching: harbors microbes.

## Slide 8: Cleaning Methods
The procedure :
Wear appropriate PPE
Remove any gross soiling on the instrument by rinsing
Take the instrument fully apart and immerse all parts in water(warm) with a detergent or enzymatic cleaner.
To prevent splashing, keep the items being washed under the surface of the water.

## Slide 9
Rinse in clean water.
Dry the instrument in a drying cabinet or by using a clean, lint-free cloth.
Inspect the instrument to ensure it is clean.
Pay particular attention to instruments with teeth, joints, or screws where organic material can collect. Open all jointed instruments
Mechanical Cleaning: Use of automated machines.

## Slide 10: High-Level Disinfection (HLD)
Removes microorganisms from instruments and equipment
Not a sterilizing process and it should not be used as a substitute for sterilization.
Used to destroy organisms on delicate or heat-sensitive instruments that cannot be sterilized .
Not appropriate for instruments that will be used in critical sites, because these instruments must be sterile

## Slide 11
Different products and processes provide different levels of disinfection:
High-level disinfection (HLD) destroys all microorganisms except some bacterial spores (especially if there is heavy contamination).
Intermediate disinfection inactivates Mycobacterium tuberculosis, vegetative bacteria, most viruses, and most fungi, but it does not always kill bacterial spores.
Low-level disinfection can kill most bacteria, some viruses, and some fungi, but it cannot be relied on to kill more resistant bacteria such as M. tuberculosis or bacterial spores.

## Slide 12
HLD is the only acceptable alternative where sterilization is not available
Can be achieved by:
Steaming: equipment/tools are disinfected by exposing them to steam from boiling water for a period of about 20mins.
Instruments are placed in a steamer pan which has holes at the bottom. This in turn is placed on top of a pan which contains the boiling water. Timing starts when steam starts getting into the instruments’ pan.

## Slide 13
Use of chemical disinfectants.
used most commonly for heat-labile equipment (for example, endoscopes) where single-use equipment is not cost-effective.
Glutaraldehyde (Cidex) is the most appropriate chemical disinfectant for high-level disinfection. However, this chemical must be used under very strictly controlled conditions and environment.

## Slide 14
Exposure can lead to:
Throat and lung irritation
Asthma and difficulty breathing
Contact and/or allergic dermatitis
Nasal irritation
Sneezing
Wheezing
Burning eyes and conjunctivitis

## Slide 15
The following are antiseptics and should not be used as disinfectants:
Acridine derivatives
Cetrimide
Chlorohexedine gluconate (hibitane)
Chloroxynelol (Dettol)
Alcohols

## Slide 16: Steps in Chemical High-Level Disinfection
Health care workers should wear PPE
Decontaminate by soaking instruments for 20 minutes in 0.5% chlorine, 2 to 4% glutaraldehyde or 6% H202.
Disassemble, clean, and dry all instruments.
Completely immerse all items in the high-level disinfectant.

## Slide 17
Remove items using HLD (or sterile) forceps and handle items wearing sterile gloves.
Rinse items well with sterile water (or boiled and filtered water) three times and air dry them.
Use items promptly or store them in a dry, HLD container with tightly fitting lid.

## Slide 18
Important
Do not use a disinfectant in an open container. There is a serious risk of contamination with multiple antibiotic-resistant bacteria such as Pseudomonas sp. and spores.
Always thoroughly decontaminate and then clean articles before disinfecting them.
When disinfectants are indicated for use on surfaces, wipe the surface. Do not wash, bathe, or flood-wash it.

## Slide 19: Sterilization
Destruction of all microorganisms (bacteria, viruses, fungi, and parasites), including bacterial endospores, from instruments and other items.
Recommended for all instruments and other items that will come in contact with the blood stream or tissues under the skin, drapes and some surgical attire.
Achieved by these methods:
High-pressure steam (autoclaving)
Dry heat (oven)
Chemical soaking(cold sterilization)

## Slide 20
Heat, either steam or dry, is the most effective method of sterilization and is reliable and cost effective if monitored carefully.
Before sterilizing any instrument or equipment ensure that it can withstand the process (e.g., steam under pressure), has been adequately cleaned, and does not require any special treatment
All materials must be wrapped before sterilization. Wrapping helps prevent recontamination after sterilization and prior to the item’s use. Only wrapped or packed sterilized materials should be described as sterile

## Slide 21: Autoclaving
Autoclaving
the use of high-pressure steam to sterilize equipment and instruments.
Types of Autoclaves:
Downward (gravity) displacement sterilizers (jacketed and nonjacketed). For sterilizing waste, solutions, and instruments.

## Slide 22
Self-contained (bench-top) sterilizers. For office-based practices. Suitable for relatively few or simple items. Do not take wrapped items; items to be used immediately .
Pre-vacuum (porous load) sterilizers. These are suitable for sterilization of clean instruments, gowns, drapes, toweling, and other dry materials that are required for surgery

## Slide 23
Autoclaving is well achieved by maintain the appropriate temperature, timing, and adequate moisture during the autoclaving cycle:
121°C throughout the process
20 minutes for unwrapped items and 30 minutes for wrapped items
100% moisture in the steam

## Slide 24: Important:
first decontaminate, clean, and dry objects
Keep instruments disassembled, opened, and unlocked.
Do not stack the instruments.
Do not wrap the packages too tightly.
Do not arrange the packs in the sterilizer too close to each other.

## Slide 25
Place bottles, solid metal, and glass containers on their sides with lids held loosely in place.
Do not overload the sterilizer or make packs too large.
Apply an autoclaving tape on the pack of instruments to indicate whether a specific temperature or pressure has been reached/achieved.

## Slide 26: Dry-Heat Sterilization Dry (hot air).
Destroys pathogens by the process of oxidation. can be achieved with a simple oven as long as a thermometer is used to verify the temperature inside the oven.
Limited value: difficult to maintain the same temperature throughout the load, while the high temperatures and the length of time required to achieve sterility make this method undesirable for many situations.
Items - withstand a temperature of 170 C (340 F).

## Slide 27: Guidelines on Dry sterilization
The door to the unit MUST not be opened while it is in the sterilizing cycle.
Wrap instruments in aluminum foil or place in a metal container with a tightly fitting, closed lid to help prevent recontamination prior to use. When using dry heat to sterilize items wrapped in cloth, be sure that the temperature does not exceed 170 C (340 F).
Instruments with cutting edges should be sterilized at lower temperatures (160 C [320 F]), because higher temperatures can destroy the sharpness of cutting edges.

## Slide 28
Place loose instruments in metal containers or on trays in the oven and heat to the proper temperature.
After the appropriate temperature is reached, begin timing.

## Slide 29
The recommended length of time depends on the temperature:
60 minutes at 170°C (340°F)
120 minutes at 160°C (320°F)
150 minutes at 150°C (300°F)
180 minutes at 140°C (285°F)
Overnight at 121°C (250°F)

## Slide 30
After cooling, remove packs or metal containers (or both) and store in a cool dry area.
Loose items should be removed with sterile forceps and used immediately or placed in a sterile container with a tightly fitting lid

## Slide 31: Chemical Sterilization
Often called cold sterilization
Alternative to high-pressure steam or dry heat sterilization, particularly for items that would get damaged by high-pressure steam or dry heat.
Glutaraldehyde is often used for chemical sterilization.
Glutaraldehyde works best at room temperature: chemical sterilization cannot be assured in cold environments (temperatures less than 20oC or 68oF), even with prolonged soaking.

## Slide 32: Guidelines on chemical sterilization
First decontaminate, clean, and thoroughly dry all instruments — water from wet instruments and other items dilutes the chemical solution, thereby reducing its effectiveness:
Prepare a 2% Glutaraldehyde solution or appropriate concentration of another chemical solution. put it in a clean container with a lid. Then mark the container with the preparation date and the expiration date.
Open and disassemble instruments accordingly.

## Slide 33
Completely submerge all instruments and other items in the solution and place any bowls and containers upright, not upside down, and fill with the solution to ensure sterilization
Sterilize for appropriate time; In 2% Glutaraldehyde, cover the container and allow the instruments and other items to soak for 10 hours. Do not add or remove any items after you begin timing.
Use large, sterile pickups (lifters, cheatle forceps) or sterile gloves to remove the instruments and other items from the solution.

## Slide 34
Rinse items thoroughly with sterile water to remove the residue that chemicals leave on instruments and other items. This residue is toxic to skin and tissues.
Place the instruments and other items on a sterile tray or in a sterile container and allow them to air dry before use. Use the instruments and other items immediately or keep them in a covered and dry sterile container and use them within one week.
Note: Sterilizing chemicals should be changed after 14 days, when they become visibly dirty, or according to manufacturers’ instructions

## Slide 35
Storage
Proper storage of sterile instruments and equipment is essential in ensuring that the product maintains its level of sterilization or disinfection till its use/utilization
Store them in a clean, dry environment that is protected from any damage.
Access to this area should be limited.

## Slide 36: Instructions for Storing Sterile Items
Keep the storage area clean, dry, dust-free, and lint-free.
Keep the temperature at approximately 24⁰C and the relative humidity below 70% when possible.
Store sacks and containers with sterile (or high-level disinfected) items 20-25 centimeters off the floor, 45-50 centimeters from the ceiling, and 15-20 centimeters from an outside wall.
Do not use cardboard boxes for storage. Cardboard boxes shed dust and debris and can harbor insects.

## Slide 37
Date and rotate the supplies (first in, first out). This process serves as a reminder, but does not guarantee the sterility of the packs.
Do not put together clean supplies, HLD items, and sterile supplies
Unwrapped items should be used immediately and should not be stored.

## Slide 38: Shelf Life
The shelf life of an item after sterilization is event related:
The item remains sterile until something causes the package or container to become contaminated—the time that has elapsed since sterilization is not always the determining factor.
To make sure items remain sterile until you need them, prevent events that can contaminate sterile packs, and protect them by placing them in protective mostly plastic covers (thick polyethylene bags).

## Slide 39
Before using any sterile item, look at the package to make sure the wrapper is clean, dry, and intact; the seal is unbroken, and no water stains are present.
If the quality of wrapping cloth is poor and plastic bags are not available, limit the shelf life to help ensure the sterility of the instruments.

## Slide 40: Central sterile services department(CSSD)
Also called sterile processing department (SPD), sterile processing, central supply department (CSD), or central supply
an integrated place in hospitals and other health care facilities that performs sterilization and other actions on medical devices, equipment and consumables; for subsequent use by health workers in the operating theatre of the hospital and for other aseptic procedures, e.g. catheterization, wound stitching and bandaging.

## Slide 41: Management of Health Care Waste
Potential reservoir of pathogenic microorganisms
Requires appropriate, safe, and reliable handling.
The purpose of proper waste management is to:
Protect people who handle waste items from accidental injury
Prevent the spread of infection to patients, clients, and HCWs

## Slide 42
Prevent the spread of infection to the local community
Safely dispose of hazardous materials (HAZMATs)
Waste from health care facility.
Noninfectious (noncontaminated-{85%}) waste poses no infectious risk to persons who handle it e.g. paper, trash, boxes, bottles.
Infectious (contaminated) waste is potentially infectious or toxic if not disposed of properly. Contaminated solid or liquid waste includes the following:

## Slide 43
Items such as sharps and used dressings that have come in contact with blood, body fluids, secretions, or excretions
Medicines, medical supplies, or other chemicals that might be toxic
Blood, body fluids, secretions, or excretions
Highly infectious e.g. body parts
Certain health care facilities may also generate hazardous waste.

## Slide 44: Principles of Waste Management
Steps in waste management :
Generation
Segregation (separation)
Collection
Transportation
Storage
Treatment
Final disposal

## Slide 45: Waste Generation
Waste Generation
Refers to the materials or products that enter a waste stream before compositing, incinerating, or recycling. Waste production.
Generated during patient management and care and in other areas of the health care setting.

## Slide 46: Waste Segregation
Waste Segregation
Separation of waste according to its level of infectivity at the source of production/generation. Divided into noninfectious, infectious, highly infectious waste and sharps
Use colour-coded bins and bin liners or label the waste containers. Use appropriate colour-coded separate containers for noninfectious, infectious, and highly infectious waste (black, yellow, red). Use sharps safety box for sharps.

## Slide 47
Waste storage
Store waste in specified areas with restricted access. Identify a storage area for waste prior to treating or moving it.
Mark the storage areas with a biohazard symbol or label it well.

## Slide 48: Waste treatment
Waste treatment
Infectious waste can be treated to render it noninfectious before disposal
Autoclave
At 121o C for a minimum of 60 minutes with heat-sensitive tape or the equivalent.